## Slide 2: Practical Terahertz Communication
A Quick Introduction to the TWISTER System and Practical DSP Concerns

## Slide 3
TWISTER
What is the TWISTER system?
Transformative Wideband continuous-wave Instrument for Science, Technology, Education, and Research

## Slide 4
The TWISTER System - Overview
It’s not an MRI machine.
Major Research Instrumentation
Keysight M8195A Arbitrary Waveform Generator
Virginia Diodes Compact Up & Down Converter Set
Two Keysight E8257D Analog Signal Generators
Keysight DSOV254A Digital Oscilloscope
Minor Research Instrumentation
Control PC
Assorted 2.92mm Cables
Antennas (horns, dishes, lenses)
Mini-Circuits ZVA-203GX+ LNA
Sketchy home-brew terahertz absorbers
Coming Soon!
Rubidium Oscillator
130 GHz wideband amplifier
High-gain dish antennas

## Slide 5
TWISTER
The Majors
Important, expensive stuff you really don’t want to break

## Slide 6
The M8195A Arbitrary Waveform Generator
It generates arbitrary waveforms!
Key Specifications
Up to 65 billion output samples per second (GSa/s) per channel.
16 GSa of waveform memory.
4 Channels*
Maximum single-channel bandwidth: 32.5 GHz
At 65 Gsa/s, the maximum waveform duration in 0.2462 seconds.

## Slide 7
Important Aside #1: Nyquist Criterion -
Nyquist Criterion
You must sample a waveform at twice the bandwidth (or faster) to preserve all information present.
You must sample at twice the highest frequency (or faster) to avoid aliasing.
The same apply when generating a signal.

## Slide 8
VDI Up and Down Converters
Up-conversion and mixing in one convenient package!
Key Takeaways
Takes in a local oscillator (LO), multiplies it to a higher harmonic, then finally mixes it with an intermediate frequency (IF) RF signal.
We have 4 VDI module pairs that collectively span 110-500 GHz.
Important Note: The VDI modules produce a double-sideband signal!

## Slide 9
Important Aside #2: Harmonics -
Generating Harmonics
Rectifying a signal generates harmonics.
In general, a nonlinear medium is needed to make harmonics.

## Slide 10
Important Aside #3: Mixing -
Important Concept
Multiplying two sinusoids together produce two sinusoids:
One at the sum frequency, and one at the difference frequency.
By generating and mixing harmonics, you can produce any integer multiple of the original sinusoid.
(Note: Your efficiency may be poor)

## Slide 11
E8257D PSG (Analog Signal Generators)
They make signals the old-fashioned way!
Key Takeaways
Output frequency from 100 kHz to 40 GHz.
Output power up to 14 dBm.
… but why?

## Slide 12
Digitization Noise
Digital waveforms aren’t necessarily clean.
Digitization noise is caused by finite digital precision.

## Slide 13
Digitization Noise
Digital waveforms aren’t necessarily clean.
Another Example

## Slide 15
The DSOV254A Digital Oscilloscope
It sees all.
Key Takeaways
It captures all the signals that we produce.
Maximum sample rate of 80 GSa/s.
  That’s 12.5 picoseconds between samples.
Memory depth is 205 MPts / channel.
Practical Note: The front-end amplifier is only 25 GHz.
Who puts male connectors on a scope!?

## Slide 16
TWISTER
The Minors
Stuff you didn’t think about, but should have.

## Slide 17
Practical Considerations
They don’t show you this on system block diagrams!
Unsung heroes
Big ‘ole box fan to keep the rack cool.  You’ll void calibration without it.
Broadband multipath absorbers.  Your antenna characterizations will be garbage without them.
Antenna systems.  Without high gain, you’ll never overcome path loss.
2.92mm cables and adapters.  Ever tried to connect two male connectors together? Doesn’t work.
The Low-Noise Amplifier (LNA).

## Slide 18
Multipath Propagation and Standing Waves
More isn’t always merrier.
Key Takeaways
The antenna beamwidth is ~11°, and within this angle, multipath interference can occur for a line-of-sight (LOS) link.
TX
RX
Practical Note: Sometimes all you need is a bag of water!
VERSUS

## Slide 19
Low-Noise Amplifier - Justification
Problem: Noise spurs on the oscilloscope.
These are intrinsic to the scope, and set the true noise floor of your measurement.

## Slide 20
Low-Noise Amplifier - Effect
The ZVA-203GX+ : A helpful amplifier with unhelpful terminals
Key Takeaways
Ask: is your noise level set by your signal, or by your instrument?  You may need an LNA.
Attack the limiting factor on your SNR.
Increased frequency resolution does not lower your noise floor!
Wait, what?

## Slide 21
Antenna Systems
Horns, Dishes, and Lenses – oh my!
Practical Issue: Lenses versus horns versus dishes
Practical Issue: Unlimited gain?

## Slide 22
TWISTER
Sending a 305 GHz Sine Wave
How hard can it be?

## Slide 23
Waveform Generation
Beware of…
Digitization Noise
Dynamic Range
DC component
Cryptic Keysight errors caused by waveform granularity requirements
Practical Note: We’ve got code that does all this for you now. I have done something useful!

## Slide 24
Setting the LO frequencies
Remember this little diagram?
Steps
Select your converters (In this case, 220-330 GHz, which has N=12)
Pick a sideband.
Set the transmitter LO. Use:
Choose: heterodyne or homodyne mixing?

## Slide 25
Heterodyne Mixing vs Homodyne Mixing
Which sideband is this one again?
Heterodyne Mixing: Mix your signal with a different center frequency. Sidebands stay split.
Homodyne Mixing: A special case where you mix with the same frequency. Combines sidebands together.
Practical Problem: You must lock the oscillators to perform Homodyne mixing. For us, that means syncing the PSGs together and waiting 45 minutes after power-up.

## Slide 26
Setting up the ‘scope
Check all the bandwidths.
Key Takeaways
Watch out for sidebands! Filter them or sample them, but don’t alias them!
Use the built-in front-end filter to cut out unwanted noise.
Make sure the front-end isn’t cutting out your signal.

## Slide 27
Communication Waveform Generation
Terms and conditions definitely apply.
Requirements
Sample rate is between 54 and 65 Gsa/s.
Integer number of carrier cycles per waveform.
Waveform length (in samples) is a multiple of 256.
Waveform length (in time) is an exact multiple of the sample period.
Note: All but the first requirement stem from the need to send frames back-to-back.
Also Note: These requirements can’t all be simultaneously met for arbitrary combinations of data length, carrier frequency, and bit rate.  Change the carrier frequency.

## Slide 28
Communication Waveform Generation
Terms and conditions definitely apply.
Requirements
Sample rate is between 54 and 65 Gsa/s.
Integer number of carrier cycles per waveform.
Waveform length (in samples) is a multiple of 256.
Waveform length (in time) is an exact multiple of the sample period.
Note: All but the first requirement stem from the need to send frames back-to-back.
Also Note: These requirements can’t all be simultaneously met for arbitrary combinations of data length, carrier frequency, and bit rate.  Change the carrier frequency.

## Slide 29
Communication Waveform Generation

## Slide 30
QUESTIONS
?